## Sample
Residual Chlorine Initial Demonstration of Capability SM 4500-Cl G
Revision 2, August 27, 2012
Enter data in blue spaces only. The spreadsheet will automatically calculate % Recovery and Relative Standard Deviation (RSD).
Analyst: |  |  |  | (Use full name)
Method No: | 4500-CL G
Analysis: | Residual Chlorine
SOP Information: |  |  |  | (i.e. SOP #, Revision # or Revision Date)
Matrix: |  |  |  | ( e.g. Drinking Water, Wastewater)
Instrument/Meter Manufacturer: |  |  |  | (i.e. Manufacturer/Brand Name/Model #)
Instrument Serial #:
Date of Analysis:
Gel Calibration Check Standards: |  |  |  | ( i.e. Manufacturer/ Brand Name/ Catalog #)
Lot # & Expiration date: | [9 empty] | Acceptance Criteria
|  | Enter your readings in the blue boxes below | [5 empty] | Average |  | Avg %
Analyte | Known Conc. * | Units | Read 1 | Read 2 | Read 3 | Read 4 | Avg | %Rec | %RSD | Recovery | %RSD
Low Residual Chlorine (STD1) | 0.23 | mg/L | 0.24 | 0.23 | 0.23 | 0.24 | 0.235 | 102.1739 | 2.4568 | 85-115 | ≤ 15
Medium Residual Chlorine (STD2) | 0.93 | mg/L | 0.91 | 0.93 | 0.93 | 0.92 | 0.9225 | 99.1935 | 1.0379 | 85-115 | ≤ 15
High Residual Chlorine (STD3) | 1.61 | mg/L | 1.62 | 1.62 | 1.63 | 1.63 | 1.625 | 100.9317 | 0.35529 | 85-115 | ≤ 15
Signature of Analyst:
Analyst Title:
Signature of Witness:
Witness Affiliation/Title:
* Refer to step 5c in the SOP.

## Blank
Residual Chlorine Initial Demonstration of Capability SM 4500-Cl G
Revision 2, August 27, 2012
Enter data in blue spaces only. The spreadsheet will automatically calculate % Recovery and Relative Standard Deviation (RSD).
Analyst: |  |  |  | (Use full name)
Method No: | 4500-CL G
Analysis: | Residual Chlorine
SOP Information: |  |  |  | (i.e. SOP #, Revision # or Revision Date)
Matrix: |  |  |  | ( e.g. Drinking Water, Wastewater)
Instrument/Meter Manufacturer: |  |  |  | (i.e. Manufacturer/Brand Name/Model #)
Instrument Serial #:
Date of Analysis:
Gel Calibration Check Standards: |  |  |  | ( i.e. Manufacturer/ Brand Name/ Catalog #)
Lot # & Expiration date: | [9 empty] | Acceptance Criteria
|  | Enter your readings in the blue boxes below | [5 empty] | Average |  | Avg %
Analyte | Known Conc. * | Units | Read 1 | Read 2 | Read 3 | Read 4 | Avg | %Rec | %RSD | Recovery | %RSD
Low Residual Chlorine (STD1) |  | mg/L |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 85-115 | ≤ 15
Medium Residual Chlorine (STD2) |  | mg/L |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 85-115 | ≤ 15
High Residual Chlorine (STD3) |  | mg/L |  |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | 85-115 | ≤ 15
Signature of Analyst:
Analyst Title:
Signature of Witness:
Witness Affiliation/Title:
* Refer to step 5c in the SOP.